## Sheet1
Instructions | 1. Please enter details per individual row for each CMDL submitted. If there were 5 CMDL submitted for a specific study, kindly complete 5 rows.
2. Write N/A for columns that are not applicable for a specific CMDL application. For example, if there are no NOD and CMDL submission went straight to approval, then write N/A for column F, G, H and I.
3. Please include exact or as much details you can provide under Column G (NOD Details) and I (Response to NOD).
4. If no response received yet from CDRRHR, please leave relevant columns blank, and add in special remarks that response from CDRRHR is pending.
5. For CMDL applications that have multiple NOD/query rounds: (A) please enter N/A for the approval date for the first entry
(B) Enter the 2nd query in a separate row (directly below the first entry). Complete column F, G, H, I, (and J, as applicable), only. Under NOD Details, please specify the round of query (e.g. Second Round, Third Round, etc).
(C) Follow the same instructions for the third and subsequent queries.
6. Under Column N (Special Remarks), please share any significant impact to the study (e.g. Sponsor dropped the study in the Philippines, recruitment was halted, multiple PDs incurred due to lack of ancillary supplies, etc).
7. Refer to the annotation in red for actual entry completion guidance
CMDL Tracker
Company Name | Study Indication | Medical Device | Submission Date | DTN Number | NOD Receipt Date | NOD Details | NOD Response Date | Response to NOD | Approval Date | Date RA Approval | Date of IL Approval | Target Site initiation Date | Special Remarks
Parexel | Gout | One use holder | 2023-07-28 | 20230728181611 | 2023-08-16 | 1. New notarized application form following/reflecting the below clarifications: - distinguish and encode the general device name and brand of device - specific intended use of the item - specific identity/reference code of the item - complete address of legal manufacturer and manufacturing site 2. CPN/R or any equivalent document issued by RA from device's country of origin (to ensure safety and efficacy of the device) 3. CT approval from REC. If not yet available, proof of application may suffice. | 18-Aug-23/ 22-Aug-23 | 1. Revised application form (signed, notarized) 2. Declaration of conformity 3. REC approval letter for some sites, proof of submission for some sites *Additional document (not inluded in NOD) required by FDAC prior accepting and forwarding compliance to CDRRHR: Revised application letter (signed, not required to be notarized) | Pending | 2023-06-08 | 2023-07-11 | 22-Aug23 (1st SIV) | Shipment of the item on hold until CMDL availability
Parexel | Gout | Blood collection needle | 2023-07-28 | 20230728181253 | 2023-08-16 | 1. New notarized application form following/reflecting the below clarifications: - corrcet classification of the device being applied - distinguish and encode the general device name and brand of device - specific intended use of the item - specific identity/reference code of the item - complete address of legal manufacturer and manufacturing site 2. CPN/R or any equivalent document issued by RA from device's country of origin (to ensure safety and efficacy of the device) 3. CT approval from REC. If not yet available, proof of application may suffice. | 18-Aug-23/ 22-Aug-23 | 1. Revised application form (signed, notarized) 2. Declaration of conformity 3. REC approval letter for some sites, proof of submission for some sites *Additional document (not inluded in NOD) required by FDAC prior accepting and forwarding compliance to CDRRHR: Revised application letter (signed, not required to be notarized) | Pending | 2023-06-08 | 2023-07-11 | 22-Aug23 (1st SIV) | Shipment of the item on hold until CMDL availability
Parexel | Non-small cell Lung Cancer | ECG machine | 2023-08-22 | 20230822102217 | [5 empty] | 2023-02-10 | 2023-03-14 | 6-Jul-23 (1st SIV) | Item required by some sites prior enrolment
Novotech | SLE | ECG Machine | 2023-09-01 | 20230803165228 | 2023-08-16 | 1) To submit a Certificate of Product Notification/Registration or any equivalent document (to ensure safety and efficacy of the device). The submitted Certificate to Foreign Government and Certificate for Exportation of Medical Products are unacceptable as equivalents of the said requirement. 2) CDRRHR also required submission of CMDL for ECG Recording Papers (which were also included in the invoice of the ECG machine) | 2023-08-22 | 1) EC Certificate for the ECG machine was submitted. 2) CMDL Application for ECG papers will be submitted on a separate application | Pending | 2023-04-17 | 2023-05-10 | 2023-09-04 | PFDA CRS has clarified through email that the medical devices must not be removed from the existing IL, and that the existing IL may still be used for the importation of medical devices while waiting for the release of the CMDL for these medical devices from CDRRHR.
Novotech | SLE | Butterfly Needles | 2023-09-01 | 20230814091037 | N/A | N/A | N/A | N/A | Pending | 2023-05-05 | 2023-05-10 | 2023-09-25
PPD | Flu Vaccine | Hypodermic Needle PrecisionGlide 25Gx25mm | 2023-07-14 | 20230717100247 | 2023-08-11 | 1) Correction on the the Notarized application forms 2) Certificate of Product registration/notification or any equivalent document issued by the regualtory agency from the device's country of origin. Certificate of compliance is insufficient 3) Packing list - needs to be signed by proper authority, invoice number should also be reflected 4) Tabulated explanation regarding consumption of total qty being applied | 2023-08-29 | **CMDL Application cancelled** Refer to Special Remarks | N/A | 2023-07-19 | 2023-07-26 | 2023-09-15 | Last 31Aug2023, PPD was just following up on the CMDL application approval via phone. The evaluator mentioned that Clinical Research Division (CRS) and CDRRHR (Medical Device Team) had a meeting and agreed that that any approved medical device on IL will no longer need a CMDL Application. PPD sent one last attempt via email with PFDA CDRRHR to clarify if the Ancillary Medical Devices and Lab Kits Supplies can be shipped to the country with only an IL and without CMDL. PFDA Replied: "As I have confirmed with my senior, yes, there is no need for a CMDL for Ancillary MDs and Laboratory Kits that were already issued an Import License from the CDRR-PRSDD-CRS. CDRRHR's CMDL is equivalent to the CDRR-PRSDD-CRS' import license as CMDL is used for one-time shipment only." as of 01Sep2023, all medical devices initially tagged for CMDL application has an approved IL and was greenlighted for shipment to the sites.
PPD | Flu Vaccine | Hypodermic Needle PrecisionGlide 25Gx38mm | 2023-07-14 | 20230717100703 | 2023-08-11 | 1) Correction on the the Notarized application forms 2) Certificate of Product registration/notification or any equivalent document issued by the regualtory agency from the device's country of origin. Certificate of compliance is insufficient 3) Packing list - needs to be signed by proper authority, invoice number should also be reflected 4) Clarification on Total Qty | 2023-08-29 | **CMDL Application cancelled** Refer to Special Remarks | N/A | 2023-07-19 | 2023-07-26 | 2023-09-15 | **Same comment as above**
PPD | Flu Vaccine | Syringe Luer-Lock 1mL | 2023-07-14 | 20230717102908 | 2023-08-11 | 1) Correction on the the Notarized application forms 2) Certificate of Product registration/notification or any equivalent document issued by the regualtory agency from the device's country of origin. Certificate of compliance is insufficient 3) Packing list - needs to be signed by proper authority, invoice number should also be reflected 4) Clarification on Total Qty | 2023-08-29 | **CMDL Application cancelled** Refer to Special Remarks | N/A | 2023-07-19 | 2023-07-26 | 2023-09-15 | **Same comment as above**
PPD | Flu Vaccine | Syringe Luer-Lok 3mL graduation 0.1mL | 2023-07-14 | 20230717103359 | 2023-08-11 | 1) Correction on the the Notarized application forms 2) Certificate of Product registration/notification or any equivalent document issued by the regualtory agency from the device's country of origin. Certificate of compliance is insufficient 3) Packing list - needs to be signed by proper authority, invoice number should also be reflected 4) Clarification on Total Qty | 2023-08-29 | **CMDL Application cancelled** Refer to Special Remarks | N/A | 2023-07-19 | 2023-07-26 | 2023-09-15 | **Same comment as above**
PPD | Flu Vaccine | Syringe Plastipak Luer-Lock 5mL | 2023-07-14 | 20230717103659 | 2023-08-11 | 1) Correction on the the Notarized application forms 2) Certificate of Product registration/notification or any equivalent document issued by the regualtory agency from the device's country of origin. Certificate of compliance is insufficient 3) Packing list - needs to be signed by proper authority, invoice number should also be reflected 4) Clarification on Total Qty | 2023-08-29 | **CMDL Application cancelled** Refer to Special Remarks | N/A | 2023-07-19 | 2023-07-26 | 2023-09-15 | **Same comment as above**
PPD | Flu Vaccine | Digital Body Thermometer 32°C to 42.9°C (F°+C°) | 2023-07-14 | 20230717104607 | 2023-08-11 | 1) Correction on the the Notarized application forms 2) Certificate of Product registration/notification or any equivalent document issued by the regualtory agency from the device's country of origin. 3) Packing list - needs to be signed by proper authority, invoice number should also be reflected 4) Justification on the quantity to be imported vs subjects refelcted on the protocol | 2023-08-29 | **CMDL Application cancelled** Refer to Special Remarks | N/A | 2023-07-19 | 2023-07-26 | 2023-09-15 | **Same comment as above**
PPD | Flu Vaccine | Needle-21G With Holder | 2023-07-14 | 20230717104922 | 2023-08-11 | 1) Correction on the the Notarized application forms 2) Certificate of Product registration/notification or any equivalent document issued by the regualtory agency from the device's country of origin. Certificate of compliance is insufficient 3) Packing list - needs to be signed by proper authority, invoice number should also be reflected 4) Tabulated explanation regarding consumption of total qty being applied | 2023-08-29 | **CMDL Application cancelled** Refer to Special Remarks | N/A | 2023-07-19 | 2023-07-26 | 2023-09-15 | **Same comment as above**
PPD | Flu Vaccine | Needle-22G Eclipse With Holder | 2023-07-14 | 20230717105219 | 2023-08-11 | 1) Correction on the the Notarized application forms 2) Certificate of Product registration/notification or any equivalent document issued by the regualtory agency from the device's country of origin. Certificate of compliance is insufficient 3) Packing list - needs to be signed by proper authority, invoice number should also be reflected 4) Tabulated explanation regarding consumption of total qty being applied | 2023-08-29 | **CMDL Application cancelled** Refer to Special Remarks | N/A | 2023-07-19 | 2023-07-26 | 2023-09-15 | **Same comment as above**
PPD | COPD | vitalograph | TBD | TBD - NOD response for submission | N/A | N/A | N/A | N/A | N/A | 2023-07-27 | 2023-08-22 | Oct/Nov2023 | 10Aug2023: CRS-NOD regarding CMDL Application for Vitalograph, ECG Machine and Electrodes and contents of the lab kits (Needle, Needle Holder Tube) was received. 15Aug2023: FDA-CRS accepted PPD's pushback that contents of the lab kits (Needle, Needle Holder Tube) are classified as IVDs. CRS confirmed that IVD products intended for research/clinical trial are not covered by A.O No. 2018-002. Thus, no certificate of Medical Device Listing shall be issued. 04Sep2023: Citing the CMDL exemption received from the PPD Flu Vaccine study, CDRRHR confirmed that Vitalograph, ECG Machine and Electrodes still requires CMDL.
PPD | COPD | ECG machine | TBD | TBD - NOD response for submission | N/A | N/A | N/A | N/A | N/A | 2023-07-27 | 2023-08-22 | Oct/Nov2023 | **Same comment as above**
PPD | COPD | ECG electrodes | TBD | TBD - NOD response for submission | N/A | N/A | N/A | N/A | N/A | 2023-07-27 | 2023-08-22 | Oct/Nov2023 | **Same comment as above**
PPD | SLE | Tubercullin Syringe | 2023-08-22 | 20230510175054 | N/A | N/A | N/A | N/A | N/A | TBD | TBD | Nov/Dec2023
IQVIA | Lupus Erythematosus | White Sensor ECG Electrodes | 2023-06-15 | 20230627155300 | 2023-08-01 | 1. Correction to the notarized form 2. Clinical trial protocol 3. FDA CTA and EC approvals 4. Packing List - signed by proper authority 5. SEC/DTI of CRO and IoR 6. Clarificaiton re. declared quantity | 2023-08-31 | Complied, but pushed back on item#4 | N/A | TBD | 2023-07-05 | 2023-09-06
Merck Sharp & Dohme (IA) LLC | Type 2 Diabetes Mellitus | Patient Weight Scale | 2023-08-30 | 20230831102102 |  |  |  |  | Pending |  | 2023-08-29 | N/A; ongoing study | CMDL Application acknowledged by FDAC on 31Aug2023; Transmitted to CDRRHR on 04Sep2023 but IL approval has been issued by CRS on 04Sep2023 as well
IQVA | Lupus Erythematosus | ECG Paper NKUK-MLSC-125 | 2023-06-15 | 20230627155512 | 2023-08-01 | 1. Correction to the notarized form 2. CPR or any equivalent doc attesting to the safety and effectiveness of the device 3. Clinical trial protocol 4. FDA CTA and EC approvals 5. Packing List - signed by proper authority 6. SEC/DTI of CRO and IoR 7. Clarificaiton re. declared quantity | 2023-08-31 | Complied, but pushed back on item#4 | N/A | TBD | 2023-07-05 | 2023-09-06
IQVIA | Lupus Erythematosus | Cardiofax®M ECG Machine | 2023-06-15 | 20230627155722 | 2023-08-01 | 1. Correction to the notarized form 2. Clinical trial protocol 3. FDA CTA adn EC approvals 4. Packing List - signed by the proper authority 5. SEC/DTI of CRO and IoR | 2023-08-31 | Complied, but pushed back on item#4 | N/A | TBD | 2023-07-05 | 2023-09-06
IQVIA | Lupus Erythematosus | 5703000012 Eppendorf 5702R Centrifuge 230v (EU) + accessories | 2023-06-15 | 20230627155939 | 2023-08-01 | 1. Correction to the notarized form 2. CPR or any equivalent doc attesting to the safety and effectiveness of the device 3. Clinical trial protocol 4. FDA CTA and EC approvals 5. Packing List - signed by proper authority 6. SEC/DTI of CRO and IoR | 2023-08-31 | Complied, but pushed back on item#4 | N/A | TBD | 2023-07-05 | 2023-09-06
IQVIA | Lupus Erythematosus | Incubator 230v (EU) with Multirack | 2023-06-15 | 20230629135740 | 2023-08-01 | 1. Correction to the notarized form 2. CPR or any equivalent doc attesting to the safety and effectiveness of the device 3. Clinical trial protocol 4. FDA CTA and EC approvals 5. Packing List - signed by proper authority 6. SEC/DTI of CRO and IoR | 2023-08-31 | Complied, but pushed back on item#4 | N/A | TBD | 2023-07-05 | 2023-09-06
IQVIA | hidradenitis suppurativa | 3M™ Tegaderm™ + Pad Film Dressing with Non-Adherent Pad | 2023-07-19 | 20230719145958 | 2023-08-11 | 1. New notarized application following/reflecting the below clarifications: a. Tick Class C as classification of the device being applied b. Clarify the correct name of the product. There is a designated box for the general device name and for the proprietary/brand name. Please treat this a critical part of the application form. c. State the specific intended use of the dressing on the space provided; 2. Certificate of Product Notification/Registration or any equivalent document issued by the regulatory agency from the device’s country of origin. This will ensure the safety and efficacy of the device. The submitted Declaration of Conformity is insufficient, this document is only applicable for Class I/Class A Medical Devices which are marketed in EU Countries; 3. The complete Research/Clinical Trial Protocol 4. Packing List showing the quantity of the product to be imported, order date, and the packing list number. This document must be duly signed by the proper authority; 5. A copy of the proof of creation of your local establishment and your local importer – SEC, DTI Registration, or LTO, whichever is available. | 2023-08-31 | Complied, but pushed back on item#4 | N/A | 2023-05-31 | 2023-07-11 | 09Oct2023 (as of 04Sep2023)
IQVIA | hidradenitis suppurativa | 3M Micropore Tape 2.5x14m | 2023-07-19 | 20230719150433 | 2023-08-11 | 1. New notarized application following/reflecting the below clarifications: - Clarify the correct name of the product. There is a designated box for the general device name and for the proprietary/brand name. Please treat this a critical part of the application form. - State the specific intended use of the device on the space provided; 2. The complete Research/Clinical Trial Protocol 3. Packing List showing the quantity of the product to be imported, order date, and the packing list number. This document must be duly signed by the proper authority; 4. A copy of the proof of creation of your local establishment and your local importer – SEC, DTI Registration, or LTO, whichever is available. | 2023-08-31 | Complied, but pushed back on item#3 | N/A | 2023-05-31 | 2023-07-11 | 09Oct2023 (as of 04Sep2023)
IQVIA | hidradenitis suppurativa | B Braun Infusomat Space Pump | 2023-07-19 | 20230719150918 | 2023-08-11 | 1. New notarized application following/reflecting the below clarifications: - Tick Class C as classification of the device being applied -State the specific intended use of the device on the space provided; -Identify if the infusion pump comes with accessories. If yes, these accessories must be declared as an additional attachment to the notarized application form 2. The complete Research/Clinical Trial Protocol 3. Packing List showing the quantity of the product to be imported, order date, and the packing list number. This document must be duly signed by the proper authority; and 4. A copy of the proof of creation of your local establishment and your local importer – SEC, DTI Registration, or LTO, whichever is available. | 2023-08-31 | complied, pushed back on #3 | NA | 2023-05-31 | 2023-07-11 | 09Oct2023 (as of 04Sep2023)
IQVIA | hidradenitis suppurativa | B Braun Infusomat® Space Line (8700036 SP Admin Sets) | 2023-07-19 | 20230719151240 | 2023-08-11 | 1. New notarized application following/reflecting the below clarifications: a. Tick Class B as classification of the device being applied b. Clarify the correct name of the product. There is a designated box for the general device name and for the proprietary/brand name. Please treat this a critical part of the application form. c. State the specific intended use of the administration set on the space provided 2. The complete Research/Clinical Trial Protocol 3. Packing List showing the quantity of the product to be imported, order date, and the packing list number. This document must be duly signed by the proper authority; and 4. A copy of the proof of creation of your local establishment and your local importer – SEC, DTI Registration, or LTO, whichever is available. | 2023-08-31 | Complied, pushed back on #3 | NA | 2023-05-31 | 2023-07-11 | 09Oct2023 (as of 04Sep2023)
IQVIA | hidradenitis suppurativa | BD Microlance™ 3 Hypodermic Needle (20G X 1") | 2023-07-19 | 20230719151636 | 2023-08-11 | 1. New notarized application following/reflecting the below clarifications: a. Clarify the correct name of the product. There is a designated box for the general device name and for the proprietary/brand name. Please treat this a critical part of the application form. b. State the specific intended use of the needle on the space provide 2. The complete Research/Clinical Trial Protocol 3. Certificate of Product Notification/Registration or any equivalent document issued by the regulatory agency from the device’s country of origin. This will ensure the safety and efficacy of the device. 4. Packing List showing the quantity of the product to be imported, order date, and the packing list number. This document must be duly signed by the proper authority; and 5. A copy of the proof of creation of your local establishment and your local importer –SEC, DTI Registration, or LTO, whichever is available. | 2023-08-31 | Complied, pushed back on #4 | NA | 2023-05-31 | 2023-07-11 | 09Oct2023 (as of 04Sep2023)
Pfizer | [5 empty] | 1) The exact use of the medical device in the study. 2) EC certificates or Declaration of conformity for each medical device. 3) Disposal process of electronic medical devices (digital thermometers, etc). 4) Signed commercial invoice to be used in the actual shipment indicating the quantity, invoice number, date of purchase. | [6 empty] | Total of 19 NODs received, only 15 were re-confirmed as ture NOD (dropping 4 that we initially flagged as requirinf CMDL ancillary itmes); Global Pfzier decided to close the recrutiment in the Philippines due to delay in the importation that would lengthen the study timeline